Repeatability and Reproducibility

Copyright Ó 1999 by Engineered Software, Inc.

Repeatability is the variability of the measurements obtained by one person while measuring the same item repeatedly.  This is also known as the inherent precision of the measurement equipment.  Consider the probability density functions shown in Figure 1.  The density functions were constructed from measurements of the thickness of a piece of metal with Gage A and Gage B.  The density functions demonstrate that Gage B is more repeatable than Gage A.

Figure 1.  Probability density functions for the thickness of 2 gages.

Reproducibility is the variability of the measurement system caused by differences in operator behavior.  Mathematically, it is the variability of the average values obtained by several operators while measuring the same item.  Figure 2 displays the probability density functions of the measurements for three operators.  The variability of the individual operators are the same, but because each operator has a different bias, the total variability of the measurement system is higher when three operators are used than when one operator is used.  Figure 3 also displays the probability density functions of the measurements for three operators using the same scale as Figure 2.  Notice that there is more difference in the means of the measurements shown in Figure 3 than those shown in Figure 2.  The reproducibility of the system shown in Figure 3 is higher than the reproducibility of the system shown in Figure 2.

Figure 2.  Reproducibility demonstration.

Figure 3.  Reproducibility demonstration.

The most commonly used method for computing repeatability and reproducibility is the Range and Average method.  The ANOVA method is more accurate, but because of the complex mathematics involved it has been shunned historically.  With desktop computers there is no excuse for not using the more accurate ANOVA method.

## Range & Average  Method

The Range & Average Method computes the total measurement system variability, and allows the total measurement system variability to be separated into repeatability, reproducibility, and part variation.

The ANOVA method, discussed in the next section, is preferred to the average range method.  The ANOVA method quantifies the interaction between repeatability and reproducibility, and is considered to be more accurate than the average and range method.

To quantify repeatability and reproducibility using average and range method, multiple parts, appraisers, and trials are required.  The recommended method is to use 10 parts, 3 appraisers and 2 trials, for a total of 60 measurements.  The measurement system repeatability is

Repeatability=5.15 {Rd2					1

where R is the average of the ranges for all appraisers and parts, and

d2 is found in Appendix A with Z = the number of parts times the number of appraisers, and W = the number of trials.

The measurement system reproducibility is

Reproducibility=5.15 {Xranged22−Repeatability2nr			2

where Xrange is the average of the difference in the average measurements between the appraiser with the highest average measurements, and the appraiser with the lowest average measurements, for all appraisers and parts,

d2 is found in Appendix A with Z = 1 and W = the number of appraisers,

n is the number of parts, and

r is the number of trials.

The measurement system repeatability and reproducibility is

R∧R=Repeatability2+Reproducibility2		3

The part variability is

VP=5.15RPd2							4

where Rp is the difference between the largest average part measurement and the smallest average part measurement, where the average is taken for all appraisers and all trials, and

d2 is found in Appendix A with Z = 1 and W = the number of parts.

The total variability, measurement system variability and part variation combined is

VT=R∧R2+VP2						5

Example 1

The thickness, in millimeters, of 10 parts have been measured by 3 operators, using the same measurement equipment.  Each operator measured each part twice, and the data is given in Table 1.

Table 1.  Range & Average method example data.

|  | Operator |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | A |  | B |  | C |  |
| Part | Trial 1 | Trial 2 | Trial 1 | Trial 2 | Trial 1 | Trial 2 |
| 1 | 65.2 | 60.1 | 62.9 | 56.3 | 71.6 | 60.6 |
| 2 | 85.8 | 86.3 | 85.7 | 80.5 | 92.0 | 87.4 |
| 3 | 100.2 | 94.8 | 100.1 | 94.5 | 107.3 | 104.4 |
| 4 | 85.0 | 95.1 | 84.8 | 90.3 | 92.3 | 94.6 |
| 5 | 54.7 | 65.8 | 51.7 | 60.0 | 58.9 | 67.2 |
| 6 | 98.7 | 90.2 | 92.7 | 87.2 | 98.9 | 93.5 |
| 7 | 94.5 | 94.5 | 91.0 | 93.4 | 95.4 | 103.3 |
| 8 | 87.2 | 82.4 | 83.9 | 78.8 | 93.0 | 85.8 |
| 9 | 82.4 | 82.2 | 80.7 | 80.3 | 87.9 | 88.1 |
| 10 | 100.2 | 104.9 | 99.7 | 103.2 | 104.3 | 111.5 |

Repeatability is computed using the average of the ranges for all appraiser and all parts.  This data is given in Table 2.

Table 2.  Example problem range calculations.

|  | Operator |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | A |  |  | B |  |  | C |  |  |
| Part | Trial 1 | Trial 2 | R | Trial 1 | Trial 2 | R | Trial 1 | Trial 2 | R |
| 1 | 65.2 | 60.1 | 5.1 | 62.9 | 56.3 | 6.6 | 71.6 | 60.6 | 11.0 |
| 2 | 85.8 | 86.3 | 0.5 | 85.7 | 80.5 | 5.2 | 92.0 | 87.4 | 4.6 |
| 3 | 100.2 | 94.8 | 5.4 | 100.1 | 94.5 | 5.6 | 107.3 | 104.4 | 2.9 |
| 4 | 85.0 | 95.1 | 10.1 | 84.8 | 90.3 | 5.5 | 92.3 | 94.6 | 2.3 |
| 5 | 54.7 | 65.8 | 11.1 | 51.7 | 60.0 | 8.3 | 58.9 | 67.2 | 8.3 |
| 6 | 98.7 | 90.2 | 8.5 | 92.7 | 87.2 | 5.5 | 98.9 | 93.5 | 5.4 |
| 7 | 94.5 | 94.5 | 0.0 | 91.0 | 93.4 | 2.4 | 95.4 | 103.3 | 7.9 |
| 8 | 87.2 | 82.4 | 4.8 | 83.9 | 78.8 | 5.1 | 93.0 | 85.8 | 7.2 |
| 9 | 82.4 | 82.2 | 0.2 | 80.7 | 80.3 | 0.4 | 87.9 | 88.1 | 0.2 |
| 10 | 100.2 | 104.9 | 4.7 | 99.7 | 103.2 | 3.5 | 104.3 | 111.5 | 7.2 |

The average of the 30 ranges, R, is 5.20.  From Appendix A, with Z = 30 (10 parts multiplied by 3 appraisers) and W = 2 (2 trials), d2 is 1.128.  The repeatability is

Repeatability=5.155.201.128=23.7

The average reading for appraiser A is 85.5, the average reading for appraiser B is 82.9, and the average reading for appraiser C is 88.9.  To compute reproducibility, the average of the range between the appraiser with the smallest average reading (appraiser B in this example) and the appraiser with the largest average reading (appraiser C in this example) is needed.  Table 3 shows this data.

Table 3.  Reproducibility example computations.

| Part | Trial | Operator B | Operator C | R |
| --- | --- | --- | --- | --- |
| 1 | 1 | 62.9 | 71.6 | 8.7 |
| 2 | 1 | 85.7 | 92.0 | 6.3 |
| 3 | 1 | 100.1 | 107.3 | 7.2 |
| 4 | 1 | 84.8 | 92.3 | 7.5 |
| 5 | 1 | 51.7 | 58.9 | 7.2 |
| 6 | 1 | 92.7 | 98.9 | 6.2 |
| 7 | 1 | 91.0 | 95.4 | 4.4 |
| 8 | 1 | 83.9 | 93.0 | 9.1 |
| 9 | 1 | 80.7 | 87.9 | 7.2 |
| 10 | 1 | 99.7 | 104.3 | 4.6 |
| 1 | 2 | 56.3 | 60.6 | 4.3 |
| 2 | 2 | 80.5 | 87.4 | 6.9 |
| 3 | 2 | 94.5 | 104.4 | 9.9 |
| 4 | 2 | 90.3 | 94.6 | 4.3 |
| 5 | 2 | 60.0 | 67.2 | 7.2 |
| 6 | 2 | 87.2 | 93.5 | 6.3 |
| 7 | 2 | 93.4 | 103.3 | 9.9 |
| 8 | 2 | 78.8 | 85.8 | 7.0 |
| 9 | 2 | 80.3 | 88.1 | 7.8 |
| 10 | 2 | 103.2 | 111.5 | 8.3 |

The average of the ranges, Xrange, is 7.015.  From Appendix A, using Z = 1 and W = 3 for 3 appraisers, is 1.91.  The reproducibility is

Reproducibility=5.157.0151.912−23.72102=18.2

The repeatability and reproducibility is

R∧R=23.72+18.22=29.9

The part variability is computed using the difference between the largest and smallest part measurement, where the average is taken for all parts and appraisers.  This data is shown in Table 4.

Table 4.  Example part variability computations.

|  |  | Operator |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  | A |  | B |  | C |  |  |  |  |
| Part | Trial 1 |  | Trial 2 | Trial 1 | Trial 2 |  | Trial 1 | Trial 2 |  | Avg |
| 1 | 65.2 |  | 60.1 | 62.9 | 56.3 |  | 71.6 | 60.6 |  | 62.78 |
| 2 | 85.8 |  | 86.3 | 85.7 | 80.5 |  | 92.0 | 87.4 |  | 86.28 |
| 3 | 100.2 |  | 94.8 | 100.1 | 94.5 |  | 107.3 | 104.4 |  | 100.22 |
| 4 | 85.0 |  | 95.1 | 84.8 | 90.3 |  | 92.3 | 94.6 |  | 90.35 |
| 5 | 54.7 |  | 65.8 | 51.7 | 60.0 |  | 58.9 | 67.2 |  | 59.72 |
| 6 | 98.7 |  | 90.2 | 92.7 | 87.2 |  | 98.9 | 93.5 |  | 93.53 |
| 7 | 94.5 |  | 94.5 | 91.0 | 93.4 |  | 95.4 | 103.3 |  | 95.35 |
| 8 | 87.2 |  | 82.4 | 83.9 | 78.8 |  | 93.0 | 85.8 |  | 85.18 |
| 9 | 82.4 |  | 82.2 | 80.7 | 80.3 |  | 87.9 | 88.1 |  | 83.60 |
| 10 | 100.2 |  | 104.9 | 99.7 | 103.2 |  | 104.3 | 111.5 |  | 103.97 |

The part with the largest average belongs to part 10, 103.97.  The lowest average belongs to part 5, 59.72.  This difference , 44.25, is Vp.  From Appendix A, using Z = 1 and W = 10 for 10 parts, d2 = 3.18.  The part variability is

VP=5.1544.253.18=71.7

The total measurement system variability is

VT=29.92+71.72=77.7

## Analysis of Variance Method

The analysis of variance method (ANOVA) is the most accurate method for quantifying repeatability and reproducibility.  In addition, the ANOVA method allows the variability of the interaction between the appraisers and the parts to be determined.

The ANOVA method for measurement assurance is the same statistical technique used to analyze the effects of different factors in designed experiments.  The ANOVA design used is a two-way, fixed effects model with replications.  The ANOVA table is shown in Table 5.

Table 5.  Two-Way ANOVA Table.

| Source of Variation | Sum of Squares | Degrees of Freedom | Mean Square | F Statistic |
| --- | --- | --- | --- | --- |
| Appraiser | SSA | a-1 | MSA=SSAa−1 | F=MSAMSE |
| Parts | SSB | b-1 | MSB=SSBb−1 | F=MSBMSE |
| Interaction (Appraiser, Parts) | SSAB | (a-1)(b-1) | MSAB=SSAB(a−1)(b−1) | F=MSABMSE |
| Gage (Error) | SSE | ab(n-1) | MSE=SSEab(n−1) |  |
| Total | TSS | N-1 |  |  |

SSA=i=1a(Yi..)2bn−Y··2N						6

SSB=j=1b(Y.j.)2an−Y··2N					7

SSAB=i=1aj=1b(Yij.)2n−Y··2N−SSA−SSB				8

TSS=i=1aj=1bk=1nYijk2−Y··2N					9

SSE=TSS−SSA−SSB−SSAB				10

a = number of appraisers,

b = number parts,

n = the number of trials, and

N = total number of readings (abn)

When conducting a study, the recommended procedure is to use 10 parts, 3 appraisers and 2 trials, for a total of 60 measurements.  The measurement system repeatability is

Repeatability=5.15MSE					11

The measurement system reproducibility is

Reproducibility=5.15MSA−MSABbn				12

The interaction between the appraisers and the parts is

I=5.15MSAB−MSEn					13

The measurement system repeatability and repeatability is

R∧R=Repeatability2+Reproducibility2+I2			14

The measurement system part variation is

VP=5.15MSB−MSABan					15

The total measurement system variation is

VT=R∧R2+VP2						16

Example 2

The thickness, in millimeters, of 10 parts have been measured by 3 operators, using the same measurement equipment.  Each operator measured each part twice, and the data is given in Table 6.

Table 6.  ANOVA method example data.

|  | Operator |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | A |  | B |  |  | C |  |  |
| Part | Trial 1 | Trial 2 |  | Trial 1 | Trial 2 |  | Trial 1 | Trial 2 |
| 1 | 65.2 | 60.1 |  | 62.9 | 56.3 |  | 71.6 | 60.6 |
| 2 | 85.8 | 86.3 |  | 85.7 | 80.5 |  | 92.0 | 87.4 |
| 3 | 100.2 | 94.8 |  | 100.1 | 94.5 |  | 107.3 | 104.4 |
| 4 | 85.0 | 95.1 |  | 84.8 | 90.3 |  | 92.3 | 94.6 |
| 5 | 54.7 | 65.8 |  | 51.7 | 60.0 |  | 58.9 | 67.2 |
| 6 | 98.7 | 90.2 |  | 92.7 | 87.2 |  | 98.9 | 93.5 |
| 7 | 94.5 | 94.5 |  | 91.0 | 93.4 |  | 95.4 | 103.3 |
| 8 | 87.2 | 82.4 |  | 83.9 | 78.8 |  | 93.0 | 85.8 |
| 9 | 82.4 | 82.2 |  | 80.7 | 80.3 |  | 87.9 | 88.1 |
| 10 | 100.2 | 104.9 |  | 99.7 | 103.2 |  | 104.3 | 111.5 |

To compute the characteristics of this measurement system, the two-way ANOVA table must be completed.  The sum of the 20 readings (10 parts multiplied by 2 trials) for appraiser A is 1710.2.  The sum of the 20 readings for appraiser B is 1657.7. The sum of the 20 readings for appraiser C is 1798.0, and the sum of all 60 readings is 5165.9.  The sum-of-squares for the appraisers is

SSA=1710.22102+1657.72102+1798.02102−5165.9260=502.5

The sum of the 6 readings for each part (3 appraisers multiplied by 2 trials) is given in Table 7 along with the square of this sum, and the square of this sum divided by 6.

Table 7.  Part sum-of-squares computations.

| Part | Sum | Sum Squared | Sum Squared/6 |
| --- | --- | --- | --- |
| 1 | 376.7 | 141,902.9 | 23,650.5 |
| 2 | 517.7 | 268,013.3 | 44,668.9 |
| 3 | 601.3 | 361,561.7 | 60,260.3 |
| 4 | 542.1 | 293,872.4 | 48,978.7 |
| 5 | 358.3 | 128,378.9 | 21,396.5 |
| 6 | 561.2 | 314,945.4 | 52,490.9 |
| 7 | 572.1 | 327,298.4 | 54,549.7 |
| 8 | 511.1 | 261,223.2 | 43,537.2 |
| 9 | 501.6 | 251,602.6 | 41,933.8 |
| 10 | 623.8 | 389,126.4 | 64,854.4 |
| Total |  |  | 456,320.9 |

The sum-of-squares for the parts is

SSB=456,320.9−5165.9260=11,545.5

The sum of the 2 trials for each combination of appraiser and part is given in Table 8 along with the square of this sum, and the square of this sum divided by 2.

Table 8.  Interaction sum-of-square computations.

| Part | Appraiser | Sum | Sum Squared | Sum Squared/2 |
| --- | --- | --- | --- | --- |
| 1 | A | 125.3 | 15,700.1 | 7,850.0 |
| 2 | A | 172.1 | 29,618.4 | 14,809.2 |
| 3 | A | 195.0 | 38,025.0 | 19,012.5 |
| 4 | A | 180.1 | 32,436.0 | 16,218.0 |
| 5 | A | 120.5 | 14,520.3 | 7,260.1 |
| 6 | A | 188.9 | 35,683.2 | 17,841.6 |
| 7 | A | 189.0 | 35,721.0 | 17,860.5 |
| 8 | A | 169.6 | 28,764.2 | 14,382.1 |
| 9 | A | 164.6 | 27,093.2 | 13,546.6 |
| 10 | A | 205.1 | 42,066.0 | 21,033.0 |
| 1 | B | 119.2 | 14,208.6 | 7,104.3 |
| 2 | B | 166.2 | 27,622.4 | 13,811.2 |
| 3 | B | 194.6 | 37,869.2 | 18,934.6 |
| 4 | B | 175.1 | 30,660.0 | 15,330.0 |
| 5 | B | 111.7 | 12,476.9 | 6,238.4 |
| 6 | B | 179.9 | 32,364.0 | 16,182.0 |
| 7 | B | 184.4 | 34,003.4 | 17,001.7 |
| 8 | B | 162.7 | 26,471.3 | 13,235.6 |
| 9 | B | 161.0 | 25,921.0 | 12,960.5 |
| 10 | B | 202.9 | 41,168.4 | 20,584.2 |
| 1 | C | 132.2 | 17,476.8 | 8,738.4 |
| 2 | C | 179.4 | 32,184.4 | 16,092.2 |
| 3 | C | 211.7 | 44,816.9 | 22,408.4 |
| 4 | C | 186.9 | 34,931.6 | 17,465.8 |
| 5 | C | 126.1 | 15,901.2 | 7,950.6 |
| 6 | C | 192.4 | 37,017.8 | 18,508.9 |
| 7 | C | 198.7 | 39,481.7 | 19,740.8 |
| 8 | C | 178.8 | 31,969.4 | 15,984.7 |
| 9 | C | 176.0 | 30,976.0 | 15,488.0 |
| 10 | C | 215.8 | 46,569.6 | 23,284.8 |
| Total |  |  |  | 456,859.0 |

The sum-of-squares for the interaction between the appraisers and the parts, SSAB, is

SSAB=456,859.0−5165.9260−502.5−11,545.5=35.6

Squaring all 60 individual reading and summing the values gives 457,405.8.  The total sum-of-squares is

TSS=457,405.8−5165.9260=12,630.4

The sum-of-squares for the gage or error is

SSE=12,630.4−502.5−11,545.5−35.6=546.8

There are 2 degrees-of-freedom for the appraisers, the number of appraisers minus one; 9 degrees-of-freedom for the parts, the number of parts minus one, 18 degrees-of-freedom for the interaction between the appraisers and the parts, the number of appraisers minus one multiplied by the number of parts minus one; 59 total degrees-of-freedom; the total number of readings minus one, and 30 degrees-of-freedom for the gage, total degrees-of-freedom minus the degrees-of-freedom for the appraisers minus the degrees-of-freedom for the parts minus the degrees-of-freedom for the interaction.  Since the mean-square-error is the sum-of-square divided by the degrees-of-freedom, the ANOVA table can be completed as shown in Table 9.

Table 9.  Example ANOVA table.

| Source of | Sum of | Degrees of | Mean | F |  |
| --- | --- | --- | --- | --- | --- |
| Variation | Squares | Freedom | Square |  | Significance |
| Appraiser | 502.5 | 2 | 251.3 | 13.8 | 0.000057 |
| Parts | 11,545.5 | 9 | 1,282.8 | 70.4 | 0.000000 |
| Interaction | 35.6 | 18 | 1.98 | 0.11 | 0.999996 |
| (Appraisers, Parts) |  |  |  |  |  |
| Gage (Error) | 546.8 | 30 | 18.2 |  |  |
| Total | 12,630.4 | 59 | 214.1 |  |  |

The significance listed in Table 9 represents the probability of Type I error.  Stated another way, if the statement is made “the appraisers are a significant source of measurement variability”, the probability of this statement being incorrect is 0.000057.  This significance is the area under the F probability density function to the right of the computed F-statistic.  This value can be found using the function FDIST(F-statistic,d1,d2) in Microsoft Excel or Lotus 123, where d1 and d2 are the appropriate degrees of freedom.

Continuing with the example, the repeatability is

Repeatability=5.1518.2=22.0

Reproducibility is

Reproducibility=5.15251.3−1.98102=18.2

The interaction between the appraisers and the parts is

I=5.151.98−18.2n=cannot be computed

Obviously the variability due to the interaction cannot be imaginary (the square root of a negative number is an imaginary number); what happened?  Each mean square is an estimate subject to sampling error.  In some cases the estimated variance will be negative or imaginary.  In these cases, the estimated variance is zero.

The repeatability and reproducibility is

R∧R=222+18.22+02=28.6

The part variation is

VP=5.151,282.8−1.9832=75.2

The total measurement system variation is

VT=28.62+75.22=80.5

# Appendix A - Values of d2

| z | w |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 |
| 1 | 1.41 | 1.91 | 2.24 | 2.48 | 2.67 | 2.83 | 2.96 | 3.08 | 3.18 | 3.27 | 3.35 | 3.42 | 3.49 | 3.55 |
| 2 | 1.28 | 1.81 | 2.15 | 2.40 | 2.60 | 2.77 | 2.91 | 3.02 | 3.13 | 3.22 | 3.30 | 3.38 | 3.45 | 3.51 |
| 3 | 1.23 | 1.77 | 2.12 | 2.38 | 2.58 | 2.75 | 2.89 | 3.01 | 3.11 | 3.21 | 3.29 | 3.37 | 3.43 | 3.50 |
| 4 | 1.21 | 1.75 | 2.11 | 2.37 | 2.57 | 2.74 | 2.88 | 3.00 | 3.10 | 3.20 | 3.28 | 3.36 | 3.43 | 3.49 |
| 5 | 1.19 | 1.74 | 2.10 | 2.36 | 2.56 | 2.78 | 2.87 | 2.99 | 3.10 | 3.19 | 3.28 | 3.36 | 3.42 | 3.49 |
| 6 | 1.18 | 1.73 | 2.09 | 2.35 | 2.56 | 2.73 | 2.87 | 2.99 | 3.10 | 3.19 | 3.27 | 3.35 | 3.42 | 3.49 |
| 7 | 1.17 | 1.73 | 2.09 | 2.35 | 2.55 | 2.72 | 2.87 | 2.99 | 3.10 | 3.19 | 3.27 | 3.35 | 3.42 | 3.48 |
| 8 | 1.17 | 1.72 | 2.08 | 2.35 | 2.55 | 2.72 | 2.87 | 2.98 | 3.09 | 3.19 | 3.27 | 3.35 | 3.42 | 3.48 |
| 9 | 1.16 | 1.72 | 2.08 | 2.34 | 2.55 | 2.72 | 2.86 | 2.98 | 3.09 | 3.19 | 3.27 | 3.35 | 3.42 | 3.48 |
| 10 | 1.16 | 1.72 | 2.08 | 2.34 | 2.55 | 2.72 | 2.86 | 2.98 | 3.09 | 3.18 | 3.27 | 3.34 | 3.42 | 3.48 |
| 11 | 1.15 | 1.71 | 2.08 | 2.34 | 2.55 | 2.72 | 2.86 | 2.98 | 3.09 | 3.18 | 3.27 | 3.34 | 3.41 | 3.48 |
| 12 | 1.15 | 1.71 | 2.07 | 2.34 | 2.55 | 2.72 | 2.85 | 2.98 | 3.09 | 3.18 | 3.27 | 3.34 | 3.41 | 3.48 |
| 13 | 1.15 | 1.71 | 2.07 | 2.34 | 2.55 | 2.71 | 2.85 | 2.98 | 3.09 | 3.18 | 3.27 | 3.34 | 3.41 | 3.48 |
| 14 | 1.15 | 1.71 | 2.07 | 2.34 | 2.54 | 2.71 | 2.85 | 2.98 | 3.09 | 3.18 | 3.27 | 3.34 | 3.41 | 3.48 |
| 15 | 1.15 | 1.71 | 2.07 | 2.34 | 2.54 | 2.71 | 2.85 | 2.98 | 3.08 | 3.18 | 3.26 | 3.34 | 3.41 | 3.48 |
| >15 | 1.128 | 1.693 | 2.059 | 2.326 | 2.534 | 2.704 | 2.847 | 2.97 | 3.078 | 3.173 | 3.258 | 3.336 | 3.407 | 3.472 |